## Slide 1: Microbiological inspection of mineral water by redox-potential measurement
Dr. Olivér Reichart
Dr. Katalin Szakmár

## Slide 2: Introduction
MicroTester as a validated method is suitable for rapid microbiological testing of mineral water, carbonated water, tank and running drinking water and other types of water. The time needed for a reliable detection of microorganisms is of key importance: in water industry the real-time (or at least as fast as possible) monitoring of the microbiological properties of the production is indispensable; in public water supply the essential basis of the epidemiological and public health measures is the fast and reliable result of the microbiological inspection. Beside the most important and most widely inspected microbiological contaminants the most relevant disturbing flora was involved to the validation process as well.

## Slide 3: Theoretical base
The energy source of the growth is the biological oxidation which results in a reduction in the environment.
This is due to the oxygen depletion and the production of reducing compounds in the nutrient medium.
A typical oxidation-reduction reaction in biological systems:
[Oxidant] + [H+] + n e- [Reductant]

## Slide 4: A typical redox curve of the microbial growth
DC: Detection Criterion
TTD: Time to Detection

## Slide 5: Microorganisms
The most frequently tested contaminant microorganisms in mineral water productions are:
      Coliforms
      Escherichia coli
      Pseudomonas aeruginosa
      Enterococcus faecalis
      Total count (22 °C and 37 °C)

## Slide 6: Method validation
    Selectivity
    Linearity
    Sensitivity
    Detection limit
    Repeatability
    Robustness

## Slide 7: Selectivity 1.
Coliforms and Acinetobacter lwoffii in BBL.
(K.o.: Klebsiella oxytoca, Ent.: Enterobacter aerogenes, Citro.: Citrobacter freundii,
E.c.: Escherichia coli, Acin.: Acinetobacter lwoffii)

## Slide 8: Selectivity 2.
Micrococcus and Enterococcus in Azid broth

## Slide 9: Selectivity 3.
Pseudomonas aeruginosa, Pseudomonas fluorescens,
E. coli and Enterococcus faecalis in Cetrimid broth.

## Slide 10: Linearity
The linear relationship between the logarithm of the cell concentration and TTD values is demonstrated by the calibration curves. From the concentrated suspensions of the test microorganisms tenfold dilution series were prepared in physiological salt solution. From the members of the dilution series the redox-potential test flasks were inoculated with 1.0 ml suspension and the TTD values were determined.

## Slide 11: Linearity
Calibration curves of Coliforms

## Slide 12: Linearity
Calibration curve of E. coli

## Slide 13: Linearity
Calibration curve of Enterococcus faecalis

## Slide 14: Linearity
Calibration curve of Pseudomonas aeruginosa

## Slide 15: Linearity
Calibration curve of total count

## Slide 16: Sensitivity
Microorganism | Broth | Regression equation | Sensitivity (min/log unit)
Citrobacter freundii | BBL | TTD (min) = 1190 - 132·lgN | 132
Klebsiella oxytoca | BBL | TTD (min) = 856 – 88·lgN | 88
Enterobacter aerogenes | BBL | TTD (min) = 774 – 81·lgN | 81
Escherichia coli | BBL | TTD (min) = 596 – 68·lgN | 68
Pseudomonas aeruginosa | Cetrimid | TTD (min) = 1440 – 155·lgN | 155
Enterococcus faecalis | Azid | TTD (min) = 836 – 92·lgN | 92
The sensitivity of the measuring method was determined as the slope of the calibration curves.

## Slide 17: Detection limit
The detection limit is 1 cell/test flask, so the system is suitable for the absence/presence tests, so considerable costs and time could be saved with more membrane filters joined together.
On the base of the calibration curves the range lasted from 1 to 7 log unit.

## Slide 18: Repeatability
The repeatability calculated from the calibration curves:
SDlgN = 0.092
SDN = 100.092 = 1.24 = 24%
which complies with the requirements of microbiological methods.

## Slide 19: Quality control tests
72 bottles tested for Coliform
Testing method of Laboratory
Membrane filtering of 3x250 ml mineral water with 1 filter. Cultivation Tergitol agar at 37 °C for 48 h. One Petri dish represents 3 bottles of mineral water.
Redox-potential measurement method
Membrane filtering of 3x250 ml mineral water with 1 filter, placing 4 membranes into 1 test flask containing BBL broth. Temperature: 37 °C. One test flask represents 12 bottles of mineral water.
Positive control: 1 ml of Citrobacter freundii suspension (lgN = 3.66)

## Slide 20: Quality control test
Results of redox-potential measurement of 72 bottles

## Slide 21: Quality control test
Bottles | 1.-12. | 13.-24. | 25.-36. | 37.-48. | 49.-60. | 61.-72.
Laboratory | negative | negative | negative | negative | negative | negative
Redox | negative | negative | negative | negative | negative | negative
Results of 72 bottles test

## Slide 22: 66 bottles tested for Coliforms
Testing method of Laboratory
Membrane filtering of 3x250 ml mineral water with 1 filter. Cultivation Tergitol agar at 37 °C for 48 h. One Petri dish represents 3 bottles of mineral water.
Redox-potential measurement method
Membrane filtering of 3x250 ml mineral water with 1 filter, placing 3 membranes into 1 test flask containing BBL broth. Temperature: 37 °C. One test flask represents 9 bottles of mineral water.
Besides the mineral water two technological water samples were tested for Coliforms
Positive control: 1 ml of Escherichia coli suspension (lgN = 6.7)

## Slide 23: Quality control test
Results of redox-potential measurement of 66 bottles

## Slide 24: Quality control test
Samples | 1.-66. Bottles | Water sample 1. | Water sample 2.
Laboratory results | negative | negative | negative
Redox method | negative | negative | negative
Results of 66 bottles test

## Slide 25: Detection time of one cell
Microbe One cell detection time (h)
Escherichia coli 11
Citrobacter freundii 23
Pseudomonas aeruginosa 24
Enterococcus faecalis 15

## Slide 26: Results of industrial tests
Microbe | All measurements (piece) | Match the standard test (%) | False positive results (%) | False negative results (%)
Escherichia coli | 942 | 99,89 | 0,11 | 0,00
Coliform | 4674 | 99,87 | 0,00 | 0,13
Enterococcus | 3000 | 99,93 | 0,00 | 0,07
Pseudomonas aeruginosa | 3372 | 99,82 | 0,06 | 0,12

## Slide 27: Advantages of the redox-potential measurement
Very simple measurement technique.
Rapid method, especially in the case of high contamination.
Applicable for every nutrient broth
Especially suitable for the evaluation of the membrane filter methods.
The test costs are less than those of the classical methods, especially in the case of zero tolerance (Coliforms, Enterococcus, Pseudomonas, etc.).